INTERFACE SPECIFICATION DOCUMENT

Shekel Security Cart

Shekel Scales (2008) Ltd.

Confidential

Revision 0.1

Version History

| Revision | Date | Description | Owner |
| --- | --- | --- | --- |
| 0.0 | 03/04/23 | Baseline-draft | Sharon H. |
| 0.1 | 10/03/24 | Update protocol for the release version. Add cs, Changed version output | Sharon H. |

Table of Contents

1	Scope	4

1.1	Introduction	4

1.2	Definitions, Acronyms and Abbreviations	4

2	Communication Parameters	4

3	REQUEST FORMAT	5

3.1	Request list	5

4	Record format	6

4.1	Weight record (W) response description	6

4.2	Zero (Z) response description.	7

4.3	Tare (T) response description.	7

4.4	Clear Tare (C) response description.	7

4.5	Reboot (B) response description.	7

4.6	Version (V) response description.	7

# Scope

## Introduction

This document describes the software interface for the SHEKEL weighing solution for security scale for cart software versions.

## Definitions, Acronyms and Abbreviations

CR 	- ASCII Carriage Return (0x0D)

LF 	- ASCII Line Feed (0x0A)

_	- SPACE indication(0x20)

# Communication Parameters

The communication interface is by serial communication via RS232 with the following parameters.

Baud rate: 9600

Data size: 8 bits

Parity: none

Stop: 1 bit

# REQUEST FORMAT

The request format is a 4-byte string, where the command byte will change according to request.

| Start | Request | Stop |  |
| --- | --- | --- | --- |
| 1B | <byte> | 1B | E |

## Request list

| Request byte | Action |
| --- | --- |
| W | Record request |
| Z | Zero request |
| T | Tare |
| C | Clear Tare |
| B | Reboot |
| V | Version |

# Record format

## Weight record (W) response description

Response String format (response to weight request):

Example:

Ascii: <CR>Z00.000G<RS>_S00.000<RS>_l<LF>

Hex: 0D 5A 30 30 2E 30 30 30 47 1E 53 30 30 2E 30 30 30 1E 6C 0A

| Start byte |  | Status 1 | Weight Record (5g) | Gross\Net Indication | Seperetor | Status 2 | Total Weight Record (2g) | Seperetor | CS Byte | Stop byte |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0D |  | Z/W/U/R/H/- | 00.000 | G\N | 1E | S\M\- | 00.000 | 1E | Calculated CS | 0A |

Weight record (5g)

Conventional weight record with 5g resolution

Total weight record

Row weight record with record resolution of 2g (accuracy depending on system usually 10g\20g) used for security scale functionality.

This record reports the total weight on the scale not affected by tare functionality.

Status 1:

W – new stable weight recoed

U – unstable weight record (scale looking for stability)

R – stable weight repeated after unstable mode

Z – the center of zero indication

‘-‘ - underload indication

H – overload indication (according to full configuration)

Status 2:

S – stable weight

M – weight in motion

‘-‘ - underload indication

CS byte

The cs byte is calculated as some of the response bytes, starting with status located in index 1 and ending with 0x1e (the 17th byte).

Example:

Original string: 0D 57 30 30 2E 32 30 30 47 1E 53 30 30 2E 31 39 38 1E 7D 0A

The calculation starts with 0x57, and ends with 0x1E, the cs calculated is 0x7D.

Cs function example:

for (c = 1; c < s.length ; c++)

ans += s[c];

## Zero (Z) response description.

The weight will be zeroed as soon as it stabilizes, with a time out of 3 seconds. There is no response to this command.

Events off failure to a zero command:

- Time out

- Out of zero range weight. (without time out)

Note: Zero range is full capacity, this is a non-regulatory software.

## Tare (T) response description.

Tare will be active only on a stable weight record, with a time out of 3 seconds. There is no response to this command.

On this command activation, when requesting weight, the weight record (5g) will indicate zero value, and the G (Gross) status will be replaced with the N (Net) status.

Events off failure to a tare command:

- Time out

- Out of full range.

## Clear Tare (C) response description.

Tare will be cleared if active.

On this command activation, when requesting weight, the weight record (5g) will indicate the gross weight value, and the N (Net) status will be replaced with the G (Gross) status.

## Reboot (B) response description.

On reboot sequence, when working with USB communication the scale will reset the USB internal PCBA driver.

The HOST (PC) needs to reestablish\refresh the USB connection to the scale.

The is no response to this command.

## Version (V) response description.

The version command is a request to the embedded software version, the response is built out of the software ID and build version.

Example:

Ascii: <CR>ID 12292.0004.0110<LF>

Hex: 0D 49 44 20 31 32 32 39 32 2E 30 30 30 34 2E 30 31 31 30 0A