MAINTENANCE INTERFACE SPECIFICATION DOCUMENT

Shekel Security Cart

Shekel Scales (2008) Ltd.

Confidential

Revision 0.0

Version History

| Revision | Date | Description | Owner |
| --- | --- | --- | --- |
| 0.0 | 24/03/24 | Baseline | Sharon H. |

Table of Contents

1	Scope	4

1.1	Introduction	4

1.2	Definitions, Acronyms and Abbreviations	4

1.3	References	4

2	Communication Parameters	4

3	REQUEST FORMAT	5

3.1	Request List	5

4	Record format	6

4.1	Info (I) request description.	6

4.2	Calibration (P) request description.	6

4.3	Next – on calibration (N) request description.	7

4.4	Load default (D) request description.	7

4.5	A2D (A) request description.	7

4.6	Set Full (F) request description.	8

5	Calibration Process	9

# Scope

## Introduction

This document describes the software interface used for manufacturing and design for the SHEKEL weighing solution for security scale for cart software versions.

## Definitions, Acronyms and Abbreviations

CR 	- ASCII Carriage Return (0x0D)

LF 	- ASCII Line Feed (0x0A)

## References

[1]  PC0034 Interface document – security scale Swan

# Communication Parameters

The communication interface is by serial communication via RS232 with the following parameters.

Baud rate: 9600

Data size: 8 bits

Parity: none

Stop: 1 bit

# REQUEST FORMAT

The request format is described in the following table. where the command byte and data string will change according to request.

| Start | Request byte | Data | Stop |  |
| --- | --- | --- | --- | --- |
| 1B | <byte> | string | 1B | E |

## Request List

| Request byte | Action |
| --- | --- |
| I | System Information |
| P | Calibration |
| N | Next - on calibration |
| D | Default |
| A | A to D |
| F | Set Full |

# Record format

## Info (I) request description.

The info request format starts with the ‘I’ byte and with no data string,

Request Example:

Ascii: 	<ESC>_i<ESC>_e

Hex: 	1B 69 1B 65

The response to the info request reports the system information data,

- Software ID and build

- Round

- Full

- Voltage AD

- Temperature AD

- Zero limits activation status

- Filter Number

Response Example:

Shekel Software, Security Scale <LF>

ID 12292, VERSION 0004, BUILD 0110<LF>

Round = Scale 5g / security 2g <LF>

Full = 01000g (programmable up to 99999g)<LF>

Voltage = 38898<LF>

Temprature = 27501<LF>

Zero Limits = Off <LF>

Filter Index = 03<LF>

## Calibration (P) request description.

The calibration process starts with the P request byte followed by the load value in the following format,

| Start | Request byte | Data | Stop |  |
| --- | --- | --- | --- | --- |
| 1B | P | Load value | 1B | E |

Request Example:

Ascii: 	<ESC>_P1000<ESC>_e

Hex:	1B 50 31 30 30 30 1B 65

Response Example:

Ascii: 	<LF> Clear the platform and press 'C'<CR>

Note: see the calibration process for the complete flow.

## Next – on calibration (N) request description.

The calibration process progress is triggered by sending the ‘NEXT’ request in the following format,

| Start | Request byte | Stop |  |
| --- | --- | --- | --- |
| 1B | N | 1B | E |

Request Example:

Ascii: 	<ESC>_N<ESC>_e

Hex:	1B 4E 1B 65

Response Example:

The response changes according to the calibration step, see the calibration process (section 5) for details.

## Load default (D) request description.

The default command returns all the memory saved values to the first flash default values. The default request is in the following format,

| Start | Request byte | Data | Stop |  |
| --- | --- | --- | --- | --- |
| 1B | D | EFAULT | 1B | E |

Request Example:

Ascii: 	<ESC>_DEFAULT<ESC>_e

Hex:	1B 44 45 46 41 55 4C 54 1B 65

Response Example:

Ascii:	<CR>Set Default<LF>

## A2D (A) request description.

The AD request format starts with the ‘A’ byte and with no data string,

Request Example:

Ascii: 	<ESC>_A<ESC>_E

Hex: 	1B 41 1B 45

The response to the AD request reports the system's live row A2D data,

Response Example:

Ascii: 	<CR>0000076444<LF>

## Set Full (F) request description.

The set full request set the maximum load defined for the system and saves it to memory. the default value is 99999g and the maximum value is 99.999kg (99999). The request is in the following format,

| Start | Request byte | Data | Stop |  |
| --- | --- | --- | --- | --- |
| 1B | F | Full value in g | 1B | E |

Request Example:

Ascii: 	<ESC>_F1000<ESC>_e

Hex:	1B 46 31 30 30 30 1B 65

Response Example:

Ascii:	<CR>Set Full: 01000g <LF>

# Calibration Process

The calibration process is built as an ASCII protocol for easy use with a serial terminal.

| Request |  | Response |
| --- | --- | --- |
| <ESC>_P1000<ESC>_e |  | <LF> Clear the platform and press 'C'<CR> |
| The user will clear the platform from any load and send the ‘NEXT’ command |  |  |
| <ESC>_N<ESC>_e |  | <LF>Zero calibration <CR> |
| At this point, zero calibration is in progress. The response is repeated *<CR> The user must not touch the platform until a message to put load is returned. <LF>Put 1000 and press 'C'<CR> At this point, the user needs to load on the platform the predefined load and send the ‘NEXT command. |  |  |
| <ESC>_N<ESC>_e |  | <LF>LOAD <CR> |
| At this point, load calibration is in progress. The response is repeated *<CR> The user must not touch the platform until a message to save is returned. <LF>Press 'C' to save <CR> At this point, the user needs to remove the load from the platform and send the ‘NEXT command to save the calibration. |  |  |
| <ESC>_N<ESC>_e |  | <LF>DONE<CR> |
| At this point, the calibration is saved, and the scale performs a reboot process. |  |  |